Animal Digital X-ray Radiography System

RV-32AL  Product configuration and parameters

1  Product configuration

| No. | Configuration Item |  | Type | Qty | Remark |
| --- | --- | --- | --- | --- | --- |
| 1 | Medical diagnostic X-ray high-frequency high-voltage generator |  | HRS-32R | 1 piece |  |
| 2 | X-ray Tube Assembly |  | LQ16(XD51-20.40/125) | 1 piece |  |
| 3 | Flat Panel Detector |  | 4343Z | 1 piece | Wired |
| 4 | Collimator |  | CRUX FR04 | 1 piece |  |
| 5 | Machine | Radiographic Table | / | 1 piece |  |
|  |  | Tube Stand | / |  |  |
| 6 | Dawei animal DR image acquisition and processing software |  | / | 1 piece |  |

2  Performance parameters

| No. | Component |  | Performance parameter |
| --- | --- | --- | --- |
| 1 | Medical diagnostic X-ray high-frequency high-voltage generator |  | Output Power: 32kW |
|  |  |  | Input Voltage: 220VAC,±10%,50Hz |
|  |  |  | kV range: 40 kV~150kV |
|  |  |  | mA range: 10 mA~400 mA |
|  |  |  | ms range: 10ms~10000 ms |
|  |  |  | mAs range: 0.1mAs~320mAs |
|  |  |  | Tech mode: kV-mA-ms；kV-mAs；kV-mAs-ms |
| 2 | X-ray Tube Assembly |  | Nominal X-ray Tube Voltage: 125kV |
|  |  |  | Nominal Focal Spot Value: 1.0mm(Small Focus)/2.0mm(Large Focus) |
|  |  |  | Nominal Anode Input Power: 20kW(Small Focus)/40kW(Large Focus) |
|  |  |  | Anode Heat Content: 140KHu |
| 3 | Flat Panel Detector |  | Image Sensor: a-Si (Amorphous Silicon) TFT |
|  |  |  | Scintillator: CsI |
|  |  |  | Effective Array: 3072 ×3072 |
|  |  |  | Pixel Size: 139 um |
|  |  |  | Effective Area (H x V): 404mm×404mm |
|  |  |  | Preview Time: 3s |
|  |  |  | Imaging Time: 5s |
|  |  |  | Spatial Resolution: 3.6 lp/mm |
|  |  |  | AD Conversion: 16 bit |
|  |  |  | Data Interface: Gigabit Ethernet |
|  |  |  | Trigger Mode: Software/Prep |
|  |  |  | Dimension (L × W × H): 474mm×474mm×22.5mm |
|  |  |  | Weight: 3.5kg |
|  |  |  | Power Consumption: 20W |
| 4 | Collimator |  | Control Mode: Manual control |
|  |  |  | Maximum Radiation: 150kV |
|  |  |  | Inherent Filtration Al. Equivalent: 1mm AL |
|  |  |  | Rectangular X-ray Field Selection Set Focus Distance, SID=100cm: Min 2cm×2cm;Max 43cm×43cm |
|  |  |  | X-Ray Field Deviation: ≤1% of SID |
|  |  |  | Light Field Indicator @ SID=100cm: ≥180lux |
|  |  |  | Light Field Indicator Precision: ≤1% of SID |
|  |  |  | Light source ON time: 30s |
|  |  |  | Maximum Radiation Leakage EN 60601-1-3: <0.5mGyh |
|  |  |  | Power Supply: 24V AC/DC,Max.4A |
|  |  |  | Dimension: 206mm×230mm×143mm |
|  |  |  | Weight: 7.0kg±0.5kg |
| 5 | Machine | Radiographic Table | Motion Control: Electromagnetic control, four-way motion |
|  |  |  | Load-bearing: ≥135kg |
|  |  |  | Height of Table from Floor: 810mm,±5% |
|  |  |  | Longitudinal floating range of tabletop: 60mm, ±5% |
|  |  |  | Lateral floating range of tabletop: 240mm, ±5% |
|  |  |  | Table Size: 1400mm × 680mm,±5% |
|  |  | Tube Stand | Overall Height: 1870mm,±5% |
|  |  |  | SID: 1000mm,±5% |
|  |  | Foot Switch | Function: Tabletop motion control |
|  |  |  | Load-bearing: ≥135kg |
| 6 | Dawei animal DR image acquisition and processing software |  | The software consists of several major modules, each performing specific functions to enable the examination workflow and provide effective diagnostic imaging data: Examination registration management:Register pet information, store it in the task list, and associate it with the protocol of pet inspection. Film examination:The film examination is performed according to the registered examination site protocol, and the position can be added or deleted at any time to adjust the film dose. Diagnostic Aids:Enhancement of DR images, providing a series of diagnostic aids such as image adjustment, cropping, marking, etc. Case management:Pet cases and examination images are stored in the database and can be browsed at any time. System interfacing:Adopting standard DICOM 3.0 communication protocol, can easily access to hospital HIS, RIS, PACS and other systems, and connect to DICOM film printer. Film Printing:Support local printing and DICOM network printing. Diagnostic Report:Reports can be set up according to the needs of hospitals or departments, and the report layout can be edited. |

- 3  Product appearance